Reversed Phase HPLC-UV Method for Biodiesel Analysis

SOP updated: Feb 3, 2024 by Griffen Gonzalez and Aaron Socha

Adapted from:  Murphy, KM. (2012) Analysis of Biodiesel Quality Using Reversed Phase High-Performance Liquid Chromatography.  Available for download here:

Scope:  Method is used after stage 1 and stage 2 reactions to quantify residual di- and triglyceride esters.  Can possibly be used for biodiesel esters and mono-glycerides but peak overlap makes this region less reliable for quantification.

Instrument:  Shimadzu HPLC

Detector:  Shimadzu DAD

Column:  Kinetex® 2.6 μm C18 100 Å LC Column 150 x 4.6 mm + SecurityGuardTM
ULTRA Cartridges UHPLC C18 4.6mm Guard column

Method:  C18_biodiesel_121823

Solvent A:  85% MeCN 15% H2O +0.1% TFA (or HOAc)

Solvent B: Acetone + 0.1% TFA (or HOAc)

| Time (min) | Solvent A (%) | Solvent B (%) |
| --- | --- | --- |
| 0 | 100 | 0 |
| 10 | 100 | 0 |
| 20 | 70 | 30 |
| 40 | 20 | 80 |
| 50 | 0 | 100 |
| 60 | 0 | 100 |

Sample concentration:  25 mg/mL in THF

Injection volume:  5 mL

Supplies Needed:

- 20 mL scintillation vial

- GC vial

- THF

- 27µL of sample

- 1000µL pipette

- 20µL pipette

- 3mL syringe

- 0.45 micron Luer lock nylon filter

I.  Biodiesel Sample Preparation:

- Tare scintillation vial and record mass.

- Using the 20µL pipette, sample from biodiesel top phase: 2 x 13.5µL aliquots (approx. 25mg of sample) into the tared scintillation vial.

- Weigh sample on analytical balance.  Subtract vial mass.  Sample mass should be approximately 25mg.

- Add calculated value (X) of THF to the scintillation vial to prepare a 25 mg/mL solution of the sample.

Sample(mg)X(µL)=25mg1000µL

- Draw the solution into a 3mL syringe and attach Luer lock nylon filter (0.45 micron)

- Filter sample into a GC vial

- Label vial with appropriate name

- Add sample to HPLC sample tray and note position

II.  Booting HPLC and choosing the Method:

*make sure all needed power (red tape) is ON before continuing*

1) Click on LCsolution software

2) Click top left most analysis labeled 1 (opening data acquisition)

3) File -> Open method -> Go back to data folder -> Orgo_Lab -> C18_biodiesel_121823

4) Simple Settings should display:

- 70.01 Time Program

- .700mL/min pump flow

- 0% solvent B, C, and D conc.

- 50ºC Temp

5) Check the volumes of solvent A and B to ensure there is enough to run all samples assuming volume requirement of each solvent as ~50 mL/sample.

6) Turn on pump/oven using icons at top bar of software.  Error light on pump can be cleared with CE button on physical pump.  It will also clear when method starts.

7) Wait for pump pressure to equilibrate (approx. 10 min using initial conditions of 100% solvent A, should be about 1030 psi)

III.   Preparing the Sample Sequence:

- Once samples are in the tray, and tray inserted until a click noise is heard, on the left of the screen select Batch Processing.

- Scroll down and add new row and copy the row with the same method being used (_biodiesel_121823.icm)

- Change the file number to match where the sample(s) placed (step 6 in Sample Prep)

- Be sure to use 5 mL injection volume

- Change the data file to appropriate name and folder *Leave all else the same*

- Highlight the samples needed to be run

- On the left side press the Batch Start button

- Make sure it has the “selection lines” correct, if not cancel action and press Batch Start again.

NOTE:  If monitoring the data in real time and baseline appears flat (not curved) after ~30 min, there is an unknown error and the sample has to be re-run.  This often happens on the first sample, but will not happen to samples in a sequence after a successful first sample.  The source of the error is not know but maybe due to a hastily equilibrated column (see Method step II. 7).

Data Analysis:

- Open Data Analysis from ‘Software Roadmap’

- Open data file from location specified in Sequence step III.5

- All samples are analyzed at 210 nm

- To integrate samples, enlarge section of interest and delete any integrals appearing as long red vertical lines.  You will need to re-draw the baseline.  Use icons for sample integration.

- To view peak area integrations, click View/peak table

Below are some representative samples and integral values collected in January 2024:

- THF blank

Peak assignment:

RT (min) 	lit (methyl ester)	area		compound

  - NA                   THF

- Distilled Biodiesel

Major peak assignments (red):

RT (min) 		area		compound

  - NA                   THF

11.229			11099679         ethyl linoleate

11.829			2410442

12.593			893806

14.336			55991

15.793			28159598 	ethyl oleate

19.656			57890

20.238			3130583

21.546			2510886

22.252			3781613

- Soybean oil (TG)

Soybean oil (enlarged TG region)

TGs

| 41.956 42.055 42.424 42.925 43.452 43.828 44.394 44.498 44.814 45.402 45.830 46.371 46.501 46.804 48.168 48.344 48.524 48.924 49.300 49.869 | 628662 404265 9194 2600506 19484 3296885 31291 14397 2805216 93646 1818738 42319 24434 753730 44745 26107 165356 151239 282473 436989 |
| --- | --- |

- Peanut oil / waste oil (post Stage 1 reaction analysis 01/25/24). Griffen G.

RT		area at 210 nm

10.261		126657

10.889		8596050 main ethyl ester

11.431		1638600

12.143		247062

12.904		90421

15.013		14773971

17.221		2434

17.856		33906

18.415		124489

19.431		42226

20.195		2550047

26.086		6998988

29.452		282781

Diglyceride section 31-39 min:

31.611		939479 * is this diglyceride

32.180		152325

33.554		160966

33.812		867522 * is this diglyceride

35.827		147818

35.932		46656

37.584		48851

37.704		55200

Triglyceride section:  43.0-48.0 min

44.105		2861282

44.411		14532

44.520		51333

44.585		61345

44.722		34075

44.990		4829910

45.369		86651

45.475		58520

45.579		52853

45.724		120043

45.940		3719717

46.180		77375

46.368		103087

46.474		188580

46.540		198975

46.880		2790038

47.237		9766

47.487		740542

47.981		22740

48.026		11734

48.061		16743

48.174		109330

48.284		27374

48.387		13979

48.463		11310

48.755		23206

48.833		47509

48.930		61369

49.052		80406

49.117		65579

49.220		83024

49.295		43438

49.379		109348

49.470		68830

49.627		76071

49.690		25304

49.941		254997

Griffen Test Waste Oil 1st Stage Failure

Triglyceride region

| Peak Number | Ret. Time | Area |
| --- | --- | --- |
| 1 | 43.179 | 638250 |
| 2 | 44.105 | 3355772 |
| 3 | 44.990 | 4764341 |
| 4 | 45.940 | 3678500 |
| 5 | 46.880 | 2906551 |
| 6 | 47.501 | 335366 |
| 7 | 47.790 | 665270 |
| Total | 4.611 |  |

Diglyceride region

| Peak Number | Ret. Time | Area |
| --- | --- | --- |
| 1 | 31.785 | 5750486 |
| 2 | 33.739 | 12700078 |
| 3 | 35.836 | 9521349 |
| 4 | 37.889 | 2362971 |
| Total |  |  |